## CAMP station data parameters.

Contents of 01_PluvA text files for 1 minute non-qc’d data
File: 01_Pluv_A: Weighing Precipitation Gauge
Order	Field	Description	Comment
3	StationID	Station Id	Station Number (stno)
1	TimeStamp	Observation data and time	End time and date of the observation period

2	Record	Record number	The Record No is generated by the datalogger automatically. This should reset when no USB connection and the power is re-cycled.
4	PluvioIntensityRT	Real Time precipitation Intensity (mm/hr)	Real-time 1 minute precipitation intensity. (mm/hr). Moving precipitation growth over the last minute before the sampling interval (measuring method according to WMO Guideline No. 8). This measured value is particularly suited for the exact determination of intensity with heavy precipitation and for alarm management. Not suited for daily/monthly totals. Threshold is 0.1mm/min – 6mm/h. / ‐99 indicates no data
5	PluvioAccRT_NRT	Uncorrected real-time 1 minute precipitation (mm)	Real-time (RT) 1 minute precipitation (mm). Values is a combination between real-time and non-real-time output. This value provides the benefit of faster RT output together with subsequent NRT output delivering the maximum accuracy possible. It shows the accumulated amounts of precipitation over the sampling period (1min). If the precipitation exceeds the thresholds, the OTT Pluvio2 L outputs the measurement result in real-time. Otherwise it collects the fine precipitation amount for 1h and outputs the measured value in non-real-time. If the fine precipitation does not reach the threshold within 1h, there will be no output. This measured values is similar the behaviour of gauge with tipping bucket. Value suited for daily and monthly totals and for alarm management. Threshold is 0.05mm within an hour. / ‐99 indicates no data
6	PluvioAccNRT	Corrected non-real-time 1 minute precipitation (mm).	Non-real-time (NRT) 1 minute precipitation (mm). / This measured value outputs the sum of the corrected amounts of precipitation over the sampling interval of 1min, with a fixed output delay of 5min.Due to better filtering, this values provides a more precise precipitation sum. Fine precipitation is collected over a maximum of one hour and output after reaching the threshold. If the fine precipitation does not reach the threshold within 1h, there will be no output. This value is particularly suited for daily and monthly totals. / Threshold is 0.05mm within an hour. / ‐99 indicates no data
7	PluvioAccTotalNRT	Total accumulated corrected non-real-time precipitation (mm).	Totalled PluvioAccNRT values. (mm) (hourly). ‐99 indicates no data. This measured value outputs the sum of the correct amounts of precipitation since the last device start (hourly) with a fixed output delay of 5min. This value is particularly suited for daily or monthly totals as well as for tracking the plausibility of the 1min RT and NRT values.
8	PluvioBucketRT	Uncorrected real-time gauge content (mm)	The value outputs the currently measured, unfiltered bucket content - real-time (RT) in mm. It corresponds to the measured weight raw data and is subjected to higher uncertainty with regards to temperature and wind impact. The measured value is suited for quick reference measurements of the weighting mechanism and for determining the current bucket level. Do not use to calculate the amount of precipitation in an external datalogger. / ‐99 indicates no data
9	PluvioBucketNRT	Corrected non-real-time gauge amount - weight (mm)	The value outputs the currently measured, filtered bucket content – non-real-time (NRT) in mm. It corresponds to the filtered weight value and is subjected to uncertainty with regards to temperature impact. The value is particularly suited for determining the content of the bucket and for calculating the evaporation behaviour. / Do not use to calculate the amount of precipitation in an external datalogger. / ‐99 indicates no data
10	PluvioLoadCellTemp	Load cell Temperature	Internal temperature of the load cell (degC). Used for compensating temperature changes. ‐99 if no data
11	PluvioHeatingStatus	Spare	Rim Heating Status (128-Not used/present). ‐99 if no data
12	PluvioStatus	Device Status	Status value (0-ok, 1-Bucket > 80% full, 4,8-restart, 16-Weight change out of range, 32-Supply voltage < 7V, 64-1024-Alarms) ‐99 if no data

Contents of 01_suitA text files for 1 minute non qc’d data
File: 01_Suit_A: Temperature, Humidity Sensors
Order	Field	Description	Comment
3	StationID	Station Id	Station Number (stno)
1	TimeStamp	Observation data and time	End time and date of the observation period

2	Record	Record number	The Record No is generated by the datalogger automatically. This should reset when no USB connection and the power is re-cycled.
4	DryA_Flag	Indicator flag for 2m air temperature	‐9 if no sensor, > 0 if problem with DryA sensor, 0 if ok
12	DryA	Air temperature (deg C)	Average of 4 x 15 second air temperature measurements. -99 if no sensor
5	GrassA_Flag	Indicator flag for grass temperature	‐9 if no sensor, > 0 if problem with GrassA sensor, 0 if ok
13	GrassA	Grass temperature (deg C)	Average of 4 x 15 second grass temperature measurements. -99 if no sensor
6	5cmA_Flag	Indicator flag for 5cm soil temperature	‐9 if no sensor, > 0 if problem with 5cmA sensor, 0 if ok
14	5cmA	5cm soil temperature (deg C)	Average of 4 x 15 second 5cm soil temperature measurements. -99 if no sensor
7	10cmA_Flag	Indicator flag for 10cm soil temperature	‐9 if no sensor, > 0 if problem with 10cmA sensor, 0 if ok
15	10cmA	10cm soil temperature (deg C)	Average of 4 x 15 second 10cm soil temperature measurements. -99 if no sensor
8	20cmA_Flag	Indicator flag for 20cm soil temperature	‐9 if no sensor, > 0 if problem with 20cmA sensor, 0 if ok
16	20cmA	20cm soil temperature (deg C)	Average of 4 x 15 second 20cm soil temperature measurements. -99 if no sensor
9	30cmA_Flag	Indicator flag for 30cm earth temperature	‐9 if no sensor, > 0 if problem with 30cmA sensor, 0 if ok
17	30cmA	30cm earth temperature (deg C)	Average of 4 x 15 second 30cm earth temperature measurements. -99 if no sensor
10	50cmA_Flag	Indicator flag for 50cm earth temperature	‐9 if no sensor, > 0 if problem with 50cmA sensor, 0 if ok
18	50cmA	50cm earth temperature (deg C)	Average of 4 x 15 second 50cm earth temperature measurements. -99 if no sensor
11	100cmA_Flag	Indicator flag for 100cm earth temperature	‐9 if no sensor, > 0 if problem with 100cmA sensor, 0 if ok
19	100cmA	100cm earth temperature (deg C)	Average of 4 x 15 second 100cm earth temperature measurements. -99 if no sensor
20	Hum_Flag	Indicator flag for relative humidity	‐9 if no sensor, > 0 if problem with Hum sensor, 0 if ok
21	Hum	Humidity (%)	Average of 4 x 15 second relative humidity measurements, ‐99 if no sensor

Contents of 01_humiA text files for 1 minute non qc’d data

File: 01_HUMI_A: Air Temperature, Humidity, Dew Point
Field	Field	Description	Comment
3	StationName	Station Id.	Station Number (stno)
1	TimeStamp	Observation data and time.	End time and date of the observation period

2	Record	Record number	The Record No is generated by the datalogger automatically. This should reset when no USB connection and the power is re-cycled.
4	AirTemp_Avg	Dew Point (deg C)	Calculated from the internal air temperature measurements of the Humidity sensor
5	Hum_AVg	Humidity (%)	Average of 4 x 15 second relative humidity measurements, ‐99 if no sensor
6	DewPnt_Avg	Dew Point (deg C)	Calculated from the internal air temperature & relative humidity measurements of the HMP155 Humidity sensor